## QCP2350 70740 cal 7-3-19
|  | Calibration Date: | 2019-07-03 |  |  |  |  | Job No.: | L-13456
|  | Model Number: | QCP2350
|  | Serial Number: | 70740
|  | Operator: | TPC |  | .
|  | Standard Lamp: | V-042(1/3/2019)
|  | Operating Voltage Range: | 6 | to | 15 | VDC (+)
Note: The QCP2350 output is a voltage that is proportional to the log of the incident irradiance.
To calculate irradiance, use this formula:
Irradiance = Calibration factor * (10^Light Signal Voltage - 10^Dark Voltage)
|  | Dry Calibration Factor: | 3.2987e+12 | quanta/cm²·sec per volt |  |  | 5.4777e-06 | µEinsteins/cm²·sec per volt
|  | Wet Calibration Factor: | 3.5432e+12 | quanta/cm²·sec per volt |  |  | 5.8837e-06 | µEinsteins/cm²·sec per volt
Sensor Test Data and Results2)
|  |  | Sensor Supply Current (Dark): | 3.44 | mA
|  |  | Supply Voltage: | 6 | Volts
|  |  | Lamp Integrated PAR Irradiance: | 9401270000000000 | quanta/cm2·sec |  | 0.015612 | µEinsteins/cm²sec
|  |  | Immersion Coefficient: | 0.931
| Nominal Filter OD | Expected Transmission | Calibrated Trans. | Sensor Voltage | Expected Voltage | Voltage % Error | Measured Trans. | Transmission Error (%) | Test Irrad. (quanta/ cm2·sec)
4556.1219
5243.7996
| No Filter | 100% | 100% | 3.455 | 3.455 | 0% | 100% | 0 | 9401324398391654
| 0.3 | 50.12% | 36.1% | 3.014 | 3.0125 | 0.04953% | 36.2% | -0.28052 | 3403460177046105
| 0.5 | 31.62% | 27.6% | 2.8985 | 2.8959 | 0.08939% | 27.74% | -0.50248 | 2607908873292870
| 1 | 10% | 9.27% | 2.4306 | 2.4221 | 0.3505% | 9.421% | -1.6066 | 8.8578e+14
| 2 | 1% | 1.11% | 1.5199 | 1.5003 | 1.288% | 1.126% | -1.4147 | 1.0591e+14
| 3 | 0.1% | 0.0534% | 0.365 | 0.18254 | 49.99% | 0.04565% | 16.988 | 4.3457e+12
| RG780 | 0% | 0% | 0.01 | 0.0071035 | 28.97% | 0.0002387% | -100 | 76836341414.8893
|  | Dark Before: | 0.0032 | Volts
|  | Light - No Filter Hldr.: | 3.455 | Volts
|  | Dark After - NFH: | 0.011007 | Volts
|  | Average Dark | 0.0071035 | Volts
Notes:
1. Annual calibration is recommended.
2) This section is for internal use and for more advanced analysis.